Progress report for the
energy spectrometer test experiment
at ESA at SLAC
Bino Maiheu
on behalf of the T474/T491 collaboration
LCWS 2007, DESY Hamburg
LCWS 2007 - DESY Hamburg, Saturday 02/06/2007 – MDI session
BPM
BPM
BPM

---
Outline
Physics motivation
Building an energy spectrometer based upon high resolution cavity BPMs
The ESA test facility at SLAC, T474/T491, experimental setup
Putting in the chicane, beam line modifications, progress since VLCWS '06
Progress on results and data analysis, understanding systematics
Outlook

---
Physics motivation, project aim
Precision physics, constrain SM parameters, eg. top threshold properties
Constrain/exclude SUSY scenarios
Impact of beam energy on measurements, luminosity spectrum dL/dE
Top quark mass scan
See e.g. talks of F. Gournaris and S. Boogert
in top QCD session...
Uncertainty on beam energy measurement
contributes directly to the uncertainty on the
ILC physics output...

---
Proposed ILC energy spectrometer
Precision measurement : dE/E 10-4
Minimal impact on beam itself : allowed emittance growth from SR
Limited space budget in BDS ~ 60 m
Minimal impact on physics data taking for e.g. calibration runs
E ~ ∫B.dl/Θ
Magnetic chicane with high resolution beam position monitors :
cavity BPMs
Max 5 mm dispersion at center chicane : determines resolution
Emittance growth determines chicane layout
Diagnostics needed:
Gain drifts : temperature
Mechanical stability : interferometer
Magnetic fields (∫B.dl ) : NMR, Hall, fluxgate magnetometers
BPM
BPM
BPM

---
T474 test experiment at ESA, SLAC
ESA comparable repetition rate, bunch charge, energy spread as ILC
Possibility to vary bunch length, energy, charge
Easy steering with feedback system
Build an energy spectrometer prototype, using a 4 magnet chicane
Goal is to demonstrate the stability of this type of energy measurement at 10-4 level, and investigate how such a magnetic chicane can be operated most efficiently at the ILC
Operate at ~5 mm ηX at center chicane as in current ILC design
Need < 1 µm resolution on position measurement (BPM)
With position measurement stability over multiple hours of ~ 100 nm

---
T474/T491 collaboration, running
Institutes involved
SLAC, UC Berkeley/LBNL, Notre Dame, Dubna, DESY/Zeuthen, RHUL, UCL, Cambridge
PI's
M. Hildreth ( Notre Dame ), Y. Kolomensky ( Berkeley ) and S. Boogert ( RHUL )
FY06 running :
January run : test run (4 days), commissioning of steering BPMs
April run :
commissioning of RF cavity BPMs outside of chicane ( old & new )
optimization of digitization and processing
July run :
commissioning of interferometer system on ILC linac prototype BPMs
commissioning of energy BPM at high dispersion
stability data taking with 10 BPMs, frequent calibrations
FY07 running :
March run :
commissioning of second energy BPM
installation and commissioning of magnetic chicane : first chicane data !
relocation of BPM/interferometer to center of chicane,
cal tone system and new processors with remotely controllable attenuation
Planned July run

---
Different BPM systems used
ILC linac prototype cavities
36 mm aperture, 2.859 GHz
low Q (~ 500)
good monopole suppression
x/y polarizations in same cavity
Middle BPM on x/y mover system
Referenced to downstream Q
Rectangular cavities, Q, x and y
Polarizations separated
2.856 GHz, high Q ~ 3000
20 mm aperture (0.8 “)
SLAC cavities
X
Y
Q
ILC cold linac prototypes

---
New spectrometer BPM prototype
Optimized design : A. Lyapin/UCL :
high resolution : ~ 100 – 200 nm
aperture
monopole suppression
own reference cavity
developed by UCL/RHUL/MSSL
mechanically rigid mover system
installation planned in July
Mover system
Use SLC MDL 2856
Digitize at 22 MHz
Dipole cavity, 2878 MHz
Processor electronics
Reference cavity
Reference cavity

---
Magnetic measurements
Simulation of magnets carried out by N. Morozov ( Dubna )
prepare for measurements in SLAC testlab (SLAC/Dubna/Zeuthen)
Main simulation results :
magnetic field integral 10-4 uniformity region is ±15 mm
region for possible NMR probe use determined ( X*Z= ±7*±40 cm)
relative contribution of the fringe field to the total field integral is 22%
maximal level of the magnetic field in return yoke is no more 0.4 T
temperature factor for the magnetic field integral is 6.1×10-5×1/C°
Screens to reduce fringe fields
Model simulation using “Vector Fields”
N. Morozov/Dubna

---
Magnetic measurements
Results of magnetic measurements in SLAC lab, Nov. '06 (SLAC/Dubna/Zeuthen)
Magnetic field integral RMS stability : 60 ppm (near working point – 150 A)
Bdl relative RMS stability : ~ 100 ppm ( both at 150 A and 200 A)
measured temperature factor for the magnetic field integral is 5.7 10-5 1/C°
in a good agreement with estimated one from magnetic field simulations 6.1 10-5 1/C°
∫B.dl value (~ 0.117 T.m when I ~ 150 A) is in agreement with simulations : 0.118 T.m
Analytical dependence of ∫B.dl vs. Current obtained in the vicinity of the working point
∫B.dl = 0.7813 10-3 x Current
Residual magnetic field along full chicane length
(vertical component)
Magnets in beam line
S. Kostromin/Dubna

---
Interferometer
Sub-nm resolution, installation itself is stable over 1 hour within 30 nm
with fixed mirrors
Relocated for march '07 run ( previously on ILC cold linac prototype triplet )
Now monitor center of chicane + one head left for new UK BPM
1 BPM in front of chicane, send laser beam down long pipe
Future location of new
UK BPM
Middle of chicane
BPM3
BPM4

---
Stability results, FY06 running
Paper “Commissioning of Spectrometer BPMs in End Station A” in the make
In depth analysis of 48 hour stability data taking with frequent calibrations
Variation in ω and γ over 48 hours found negligible ( ~ 2 nm )
justifies use fixed ω and γ in algorithm which extracts amplitude and phase
Calibration with beam ( corrector scans ) induce quite large variations in
IQ phase and scales : no feedback, beam jitter
mover calibrations + think of clever calibration scheme (helmholtz coils)
Commissioned during march '07 run
Fast pulse per pulse beam motion
Less sensitive to beam jitter/drifts
Average scales and IQ phases
Automation : write status into ADC

---
Stability results, FY06 running
X
Y
Using SVD algorithm to predict
position from spectator BPMs
Resolution given by RMS of residual
Stability given by drift of residual
Typical good Resolution :
~ 700 nm for ILC cavities
~ 300 nm for SLAC cavities
BPM3
+/- 100 nm
SLAC cavities
M. Slater/Cambridge
ILC cold linac prototype
X
Y
Drifts can be seen over course of 1 hour
no gain monitoring available yet
correlate with temperature
Long term drifts : clearly degradation of
resolution
Orbit stability shows to be +/- 1 μm over
the course of an hour and RMS on the
level of 1 μm
Need to understand systematics way
better : calibration tone, more analysis !!!
BPM10
1 hour

---
Calibration tone data, March '07 run
Send triggered CW tone down electronics and monitor gain / phase drifts
UCL/RHUL/UC Berkeley
Correlate with temperature
Calibration tone test runs March '06 (e.g. BPM 41/42)
Temperature monitoring on processors
x42Amp/q42Amp
x42Phase
M.Christiakova/Berkeley
Clear correlation, phase seems fairly stable, need to analyze what effect this has on
stability of energy measurements, data needed !

---
Preliminary spectrometer results
Taking into account ∫B.dl and deflection at center of chicane, can
compute correct beam energy
Have to subtract incoming orbit in each event : prove we measure just energy !
Further detailed analysis, spectrometer stability studies underway...
More and better data to come in July...
S. Kostromin/Dubna
PRELIMINARY !!

---
Outlook
Planned July run
Installation of new BPM prototype at center of chicane on mover system
Further data taking, full commissioning of calibration system
will be first good run with complete chicane up and running
also commissioning of synch stripe measurements (T475) :
additional energy measurement systematics !
Future plans to install metrology grid (M. Hildreth et al.)
Understanding mechanical stability
Metrology Grid : Crucial for Mechanical Stability Tests
“straightness monitor”
~1m
“reference bar” for triangulation. Contains internal interferometer for monitoring (LiCas)
Publications in pipeline
More systematic understanding of complete system needed !
BPM Station
BPM Station
Interferometer arms

---
Backup Slides

---
Stability/resolution over 48 hours

---
Linked resolution

---
Helmholtz calibration results

---
Helmholtz calibration results